## Slide 1: ANEUVAS TECH. INC. PORTABLE MEDICAL BENCH
DR. BECKER - ADVISOR and CLIENT
Kenyon Rowley | Project Manager and Financial Manager
Katherine Riffle | Test Engineer and CAD Engineer
Hunter Daniel | Logistics Manager and Manufacturing Engineer

## Slide 2: Project Description
Safely and easily transporting blood flow model experimental setup through buildings
Compatible and Protective of Medical Research Devices
Support Clean-Room Hood
Reduce Shock During Transport
X-Ray Machine Compatible
Spill Prevention
Fig 1: X-Ray Machine
Fig 2: Clean Room Hood and Air Filter
Kenyon Rowley, April 23, 19F05 - Aneuvas Tech Portable Bench
2

## Slide 3: Specifications
Engineering Requirement | Units | Target Value | Tolerance | Met?
Cost | $ | 1000 | +/- 100 | Approved
Weight | lb | 150 | +/- 10 | Test
Fitting Through Doorway | ft2 | 7.5 | +/- .1 | Test
Tabletop Yield Strength | psi | 5 | +/- 1 | Test
Effective shock absorption | in/s2 | 5 | +/- 5 | Approved
Tabletop Deflection | in | 0.25 | +/- .05 | Test
Tabletop Thickness | in | 1.00 | +/- 0.10 | Met
Bench Height | in | 36.00 | +/- 0.10 | Met
Storage Volume | ft3 | 5 | +/- 1 | Met
Temperature resistance | °F | 50 | +/- 50 | Test
Liquid Drained | % | 80 | +/- 10 | Test
Table 1: Engineering Requirements
Kenyon Rowley, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 4: Design Solution
Table Top material selected was spill resistant and x-ray compatible
Table top, frame, and storage underwent multiple versions to ensure the safe and efficient transportation of equipment
Figure 3: Final CAD
Kenyon Rowley, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 5
Design Description: Tabletop
Figure 4: Tabletop with Wedges CAD
Spill Guards attached around hood perimeter
No grooves for hood
Wilsonart - no polycarbonate workspace
No taper, only frame angled
Wedge attachments for frame
Kenyon Rowley, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 6
Design Description: Frame
Bolt Plates for Wheel Attachment, 1” Tilt from back to front
Air Filter Support Shelf for Transport
Figure 5: Frame CAD and Constructed
Katherine Riffle, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 7
Design Description: Wheels
Fig 6: Wheels, CAD and ordered
Selected for weight capacity and shock absorption
Bolted onto frame
Katherine Riffle, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 8
Design Description: Storage
Fig 7: Storage CAD
Stores the 40lbs of devices
Two of three devices stored underneath: Reservoir, Super pump, and controller
Drawer to hold pens, pencils, etc. for multipurpose use
Katherine Riffle, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 9: Decision Processes
Tabletop:
Oak and Polycarbonate → Wilsonart
Tilted Table Top → Angled Frame, Wedges
Spill Guards around Perimeter of Hood
Gutter Tray
Frame:
Dimensions Match Hood → Wider
Storage for Air Filter during transport
Storage:
Storage will be removable
Drawer slides have been omitted
Fig 8: Current SolidWorks Design
Katherine Riffle, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 10: Manufacturing
Completed with a mixture of in-house and outsourced resources
Table top: Only Table Tops
Wheels: McMaster-Carr
Storage: In-house
Frame: Materials ordered, welding done in-house
Fig 9: Completed table top
Katherine Riffle, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 11: Testing
Team was unable to complete desired testing
The team was planning on completing testing in the Wettaw
These tests would have inspired design improvements
Figure 10: Completed storage
Hunter Daniel, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 12
Testing Procedures
TP |  | ER |  | Requirements/Procedure
1 | Project Cost | 1 | Cost Target | Bill of Materials Approval from Dr. Becker and Dr. Oman
2 | Bench Specifications | 7 | Tabletop Thickness | Measure the Tabletop Thickness and Bench Height
 |  | 8 | Bench Height | 
 |  | 9 | Storage Volume | Fit devices in the storage area with client supervision
3 | Transporting Bench | 3 | Fitting Through Doorway | Bring the tabletop through the Wettaw building with client escort
 |  | 5 | Shock Absorption | 
Table 2: Testing Procedures 1-3
Hunter Daniel, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 13
Testing Procedures
TP |  | ER |  | Requirements/Procedure
4 | Temperature Resistance | 10 | Temperature Resistance | Hot Plate and Generator; Cold Flagstaff Day Clean Room Hood and Air Filter for Deflection and Yield Strength Measurements during extreme temperatures
5 | Liquid Drainage/ Deterioration | 11 | Liquid Drained | Measure amount of maximum spilled liquid retained and ensure tabletop does not deteriorate from lab fluids
6 | Weight | 2 | Overall Weight | Industrial scale; client’s approval of weight
7 | Tabletop Strength | 4 | Tabletop Yield Strength | Cleanroom Hood and Air Filter placed on top
 |  | 6 | Tabletop Deflection | 
Table 3: Testing Procedures 4-7
Hunter Daniel, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 14: Budget
Hunter Daniel, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 15: Future Work
Completing the testing procedures
Evaluate how the bench completes tests
After procedures, make design changes to meet customer requirements
Hunter Daniel, April 23, 19F05 - Aneuvas Tech Portable Bench

## Slide 16: ANEUVAS TECH. INC. PORTABLE MEDICAL BENCH
DR. BECKER - ADVISOR and CLIENT
Kenyon Rowley | kjr293@nau.edu
Katherine Riffle | krr264@nau.edu
Hunter Daniel | hbd6@nau.edu

## Slide 17: Appendix A: CAD Drawings

## Slide 18: Appendix A: CAD Drawings

## Slide 19: Appendix A: CAD Drawings

## Slide 20: Appendix A: CAD Drawings

## Slide 21: Appendix B: Full Bill of Materials